|  |  |  |  |  | ‘DOCUMENT CHANGE REQUEST |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # TO BE COMPLETED BY ORIGINATOR |  |  |  |  |  |  |  |  |  | Change request No. (4) |
| Originator (1) | Originator signature (2) |  |  |  |  |  | ## NSA or ESA representative signature (3) |  |  |  |
| Jude Neylon |  |  |  |  |  |  |  |  |  |  |
| Affiliation | Date:26th October 2007 |  |  |  |  |  | Date: 7 November 2007 |  |  | Page 1 of [4] (5) |
| Betatherm Ireland Ltd. |  |  |  |  |  |  |  |  |  |  |
| ### DOCUMENT AFFECTED |  |  |  |  |  |  |  |  | Other documents affected (10) |  |
| Doc. No. (6) |  | Status (7) |  |  |  | Title (8) |  |  |  |  |
| ‘4 4006/014 |  | ‘Issue 5 July 2007 |  |  |  | THERMISTORS (THERMALLY SENSITIVE RESISTORS), NCT, RANGE 2 000 TO 15 000 OHMS AT + 25OC WITH A TEMPERATURE RANGE OF – 60 TO + 160 OC |  |  |  |  |
| Paragraph(s) and page(s) affected (9) |  |  |  |  |  |  |  |  | NONE |  |
| Pages: 1,5,6,7,8,12, 17 |  |  |  |  |  |  |  |  |  |  |
| Paragraphs: Specification Title; para.1.1; Table 1(a); Note 3 to Table 1(b); Table of Figure 2; Table 2; Appendix A. |  |  |  |  |  |  |  |  |  |  |
| ## ‘PROPOSED WORDING OF CHANGE (11) |  |  |  |  |  |  |  |  |  |  |
| SEE ATTACHED PAGES |  |  |  |  |  |  |  |  |  |  |
| x |  |  |  |  |  |  |  |  |  |  |
| Continuation sheet(s) attached: Yes No |  |  |  |  |  |  |  |  |  |  |
| ‘JUSTIFICATION (12) |  |  |  |  |  |  |  |  |  |  |
| SEE ATTACHED PAGES |  |  |  |  |  |  |  |  |  |  |
| x |  |  |  |  |  |  |  |  |  |  |
| Continuation sheet(s) attached: Yes No |  |  |  |  |  |  |  |  |  |  |
| ‘Changes required for: Procurement (project) Qualification MRB decision (13) ## X |  |  |  |  |  |  |  |  |  |  |
| A General Improvement of Spec. Other |  |  |  |  |  |  |  |  |  |  |
| #### ‘RESERVED FOR USE BY SCC SECRETARIAT |  |  |  |  |  |  |  |  |  |  |
| ## Date of registration |  |  |  | Order of priority for Appr. / Impl.: 1 (high) 2 (medium) 3 (low) |  |  |  |  |  |  |
| ## Attachments |  |  |  | Qualification status: Qualified In process of qualification |  |  |  |  |  |  |
| #### ‘RESERVED FOR USE BY APPROVING AUTHORITY |  |  |  |  |  |  |  |  |  |  |
| ## Approved |  |  | ## Date and signature |  |  |  |  | ## Reference to SCCG decision |  |  |
| Yes No |  |  |  |  |  |  |  |  |  |  |
| Priority |  |  |  |  |  |  |  |  |  |  |
| ## ‘Approved wording, if different from box 11 or reason for rejection (14) |  |  |  |  |  |  |  |  |  |  |
| Continuation sheet(s) attached: Yes No |  |  |  |  |  |  |  |  |  |  |

|  | ‘DOCUMENT CHANGE REQUEST |  |
| --- | --- | --- |
| # CONTINUATION SHEET FOR BOX [ ] |  | Change request No.) |
|  |  | Page 2 of [4] |
| ‘ |  |  |
| CHANGE: Page 1 - Title from “…RANGE 2000 TO 15000 OHMS AT + 25°C WITH…” to “… RANGE 2000 TO 100000 OHMS AT + 25°C WITH…” from “…BASED ON TYPE G15K4D489, G10K4D453 , G2K7D411 , G4K7D421” to “…BASED ON TYPE G15K4D489, G10K4D453 , G2K7D411 , G4K7D421 , G100K6D487” JUSTIFICATION: The new variant 12 (based on Type G100K6D487) has been developed to meet current customer demand. The construction is similar to the previous ESCC 400601405B Issue 4, July 2005. The main difference is that the glass bead coating and wire type / ferrule are the same as the qualified part ESCC 400601408B Issue 5, July 2006. (G15K4D489). |  |  |
| CHANGE: Page 5 Paragraph 1.1 |  |  |
| from “…Range 2000 to 15000 OHMS at + 25°C with…” to “…Range 2000 to 100000 OHMS at + 25°C with…” |  |  |
| from “ …based on type G15K4D489, G10K4D453 , G2K7D411 , G4K7D421.” to “…based on type G15K4D489, G10K4D453 , G2K7D411 , G4K7D421 , G100K6D487.” |  |  |
| JUSTIFICATION: New variant 12 has been developed from old customer specifications to meet demand. Construction is similar to variant 05, except it has different glass and lead size. |  |  |
| The new variant 12 (based on Type G100K6D487) has been developed to meet current customer demand. The construction is similar to the previous ESCC 400601405B Issue 4, July 2005. The main difference is that the glass bead coating and wire type / ferrule are the same as the qualified part ESCC 400601408B Issue 5, July 2006. (G15K4D489). CHANGE: Page 6 Table 1(a) Add variant 12 to Table, giving resistance values and appropriate resistance tolerances at temperatures +25°C; +100°C; +125°C; + 140°C:+160°C. The insertion is as follows: VARIANT; BASED ON TYPE; RZ; RESISTANCE/TEMPERATURE CHARACTERISTICS (NOTE 3). -60 OC; -40 OC; -20 OC; 0 OC; +25OC; +50OC; +70OC; +100OC; +125OC; +140OC; +160OC. 12; G100K6D487; NOM (); --; --; --; ---; 100000; --; ---; 5574; 2642.4; 1756.3; 1059.0. TOL (+%); 1.75; 1.11; 1.41; 1.32; 1.21. JUSTIFICATION: The new variant 12 (based on Type G100K6D487) has been developed to meet current customer demand. The construction is similar to the previous ESCC 400601405B Issue 4, July 2005. The main difference is that the glass bead coating and wire type / ferrule are the same as the qualified part ESCC 400601408B Issue 5, July 2006. (G15K4D489). The tolerances from +100°C to +160°C are achievable as the Betatherm ceramic curve for this product is equivalent to the Fenwal curve on which these specification tolerances were based. |

|  | ‘DOCUMENT CHANGE REQUEST |  |
| --- | --- | --- |
| # CONTINUATION SHEET FOR BOX [ ] |  | Change request No.) |
|  |  | Page 3 of [4] |
| ‘ |  |  |
| CHANGE: Page 7 Table 1(b) Change Note 3 from “ – 40OC for Variants 10, 11 and – 60OC for Variant 08 and 09 to the Maximum Operating Temperature specified in Column 4 of Table 1(a)” to “ – 40OC for Variants 10, 11, 12; and – 60OC for Variants 08 and 09 to the Maximum Operating Temperature specified in Column 4 of Table 1(a)” JUSTIFICATION: To include the storage temperature range of the new Variant 12. CHANGE: Change Page 8,Table of Figure 2, from |  |  |
| SYMBOL; MILLIMETRES. VARIANT 08; VARIANT , 09; VARIANT 10,11. MIN; MAX; MIN; MAX; MIN; MAX. A; 356; 406; 500; 550; 280; 330. B; 6.1; 6.6; 6.1; 6.6; 6.1; 6.6. C; -; 2.4; -; 2.8; -; 2.8. D; -; 9.8; -; 9.8; -; 9.8. E; 0.33; 0.48; 0.33; 0.48; 0.33; 0.48. F; -; 50; -; 50; -; 50. G; 50; 80; 50; 80; 50; 80. |
| to |  |  |
| SYMBOL; MILLIMETRES. VARIANT 08; VARIANT , 09; VARIANT 10,11,12. MIN; MAX; MIN; MAX; MIN; MAX. A; 356; 406; 500; 550; 280; 330. B; 6.1; 6.6; 6.1; 6.6; 6.1; 6.6. C; -; 2.4; -; 2.8; -; 2.8. D; -; 9.8; -; 9.8; -; 9.8. E; 0.33; 0.48; 0.33; 0.48; 0.33; 0.48. F; -; 50; -; 50; -; 50. G; 50; 80; 50; 80; 50; 80. |
| JUSTIFICATION: To include physical dimensions for the new Variant 12. |  |  |

|  | ‘DOCUMENT CHANGE REQUEST |  |
| --- | --- | --- |
| # CONTINUATION SHEET FOR BOX [ ] |  | Change request No.) |
|  |  | Page 4 of [4] |
| ‘ |  |  |
| CHANGE: Page 12, Table 2, No. 3 change from 3; Thermal Time Constant Variant 08 Variants 09,10,11; KH; Para. 9.3.1.3 Tamb = +25 + 10C In Still Air Note 3; - -; 25 40; sec.. to 3; Thermal Time Constant Variants 08 Variants 09,10,11,12; KH; Para. 9.3.1.3 Tamb = +25 + 10C In Still Air Note 3; - -; 25 40; sec.. JUSTIFICATION: To include the Thermal Time Constant characteristic of the new Variant 12. |
| CHANGE: Page 17, Appendix ‘A’; AGREED DEVIATIONS FOR BETATHERM (IRELAND), for all three “ITEMS AFFECTED” – Para. 9.2 Thermal Shock; Para. 7.4/7.4.1 Check for Lot Failure/Lot Failure during 100% Testing; and Para. 9.14.2 Operating Life during Lot Acceptance Testing. change from “For Variants 08, 09, 10, 11…….” to “For Variants 08, 09, 10, 11, 12…….” JUSTIFICATION: To apply the current agreed deviations for Betatherm Ireland to the new Variant 12. |  |  |